# HIGH PRECISION WEIGHING SCALE

## Operation Manual

● ZX-EWSB2003-03

CONTENTS

PRECAUTION BEFORE USING THE SCALE------------------------1

INTRODUCTION--------------------------------------------------------------2

COMMON SPECIFICATIONS----------------------------------------------3

DISPLAYS---------------------------------------------------------------------- 4

KEY DESCRIPTIONS--------------------------------------------------------5

OPERATION

ZERONING THE DISPLAY--------------------------------------------------7

TARING--------------------------------------------------------------------------7

WEIGHING A SAMPLE------------------------------------------------------8

PERCENT WEIGHING------------------------------------------------------8

PARTS COUNTING----------------------------------------------------------8

CHECK-WEIGHING----------------------------------------------------------9

ACCUMULATED TOTAL---------------------------------------------------10

MANUAL ACCUMULATION-----------------------------------------------11

PARAMETERS---------------------------------------------------------------12

BATTERY OPERATION----------------------------------------------------14

RS-232 OUTPUT------------------------------------------------------------15

AUTO CALIBRATION-------------------------------------------------------17

ERROR CODES-------------------------------------------------------------18

PRECAUTION BEFORE USING THE SCALE

### Environment

The scale should always be used in an environment, which is free from excessive air currents, corrosives, vibration, and temperature or humidity extremes. These factors will affect displayed weight reading.

DO NOT install the scale:

*Next to open windows or doors causing drafts or rapid

temperature changes

*Near air conditioning or heating vents

*Near vibrating, rotating or reciprocating equipment

*Near magnetic fields or equipment that generates magnetic fields

*On an unstable work surface

*In a dusty environment

*In direct sunlight

Leveling the scale

The scale is equipped with a level indicator on the left bottom side of the front panel and four adjustable leveling feet. Adjust the leveling feet until the bubble appears in the center circle of the indicator.

### Turn on the scale

Please remove shipping protection screw from the bottom of the scale before turning on!

Do not turn on scale with anything on the platform. The switch is located on the right side of the scale. The scale will start to count down from nine to zero. The scale is then ready for use, Give a warm-up for 15~30 minutes before use.

INTRODUCTION

The precision electronic weighing scale offers a range of capacities from 1.5 kilograms to 30kg kilograms.

The scales are very simple to use and applicable for general weighing. The user can also use the parts counting and percent weighing functions for special applications. Special functions are available for weighing in up to 5 different units of weight. And the scale use kilograms as the default unit.

The standard RS-232 interface allows the data to be sent to a computer or printer. (Only for the scale with RS-232 interface)

The scale is with large easy to read LCD display with backlight, and all the keypads are sealed membrane switches.

The scale include audible alarm for pre-set weights, tare, pre-set tare and an accumulation facility that allows the count to be stored and recalled as an accumulated total.

COMMON SPECIFICATIONS

| Interface | RS-232 Output Optional |
| --- | --- |
| Operating Temperature | 0℃~40℃ |
| Power supply | 12V DC, 800mA from external adapter |
| Calibration | Automatic External |
| Display | 6 digits LCD digital display |
| Pan Size | 300×230mm |
| Functions | Weighing, parts counting, % weight, check weighing |
| Other Features and Specs | Internal rechargeable battery |

DISPLAYS

The LCD display will show a value and a unit to the right of the digits.

Other labels are TARE, GROSS weight, ZERO, stable “  ” and for Low battery“       ”

#### KEY DESCRIPTIONS

### Zero

Set the zero point for all subsequent weighing. The display shows zero.

A secondary function“  ”, or “Enter” key when setting parameters or other functions.

### Tare

Tares the scale. Stores the current weight in memory as a tare value, subtracts the tare value from the weight and shows the results. This is the net weight.

A secondary function“▲”,of incrementing the active digit when setting a value for parameters or other functions.

### Check-W

Sets the limits for check weighing. Allows setting of either the low limit or the high limit or both.

Secondary function “”, is to move the active digit to the right when setting values for parameters or other functions.

%

Enters the percent weighing function. Allows the weight, unit weight, and count to be seen when parts counting.

Secondary function “”,is to move the active digit to the left when setting values for parameters or other functions.

### Func

Used to select the function of the scale. If the scale is weighing, it will select parts counting. Of it is not in weighing mode it will return the user to weighing.

Secondary function (c), is to act as a clear key when setting values for parameters or other functions.

### Print

To print the results to a PC or printer using the optional RS-232 interface. It also adds the value to the accumulation memory if the accumulation function is not automatic.

Secondary function (ESC), is to return to normal operation when the scale is in a parameter setting mode.

### Units

Select weigh units (kg, g , oz, lb, t)

### OPERATION

### ZERONING THE DISPLAY

You can press the ZERO/ENTER key at any time to set the zero point from which all other weighing and counting is measured, within 4% of power up zero.  This will usually only be necessary when the platform is empty. When the zero point is obtained the display will show the indicator for zero.

The scale has an automatic re-zeroing function to account for minor drifting or accumulation of material on the platform. However you may need to press the ZERO/ENTER key to re-zero the scale if small amounts of weight are shown when the platform is empty.

### TARING

Zero the scale by pressing the ZERO/ENTER key if necessary. The zero indicator will be on.

Place a container on the platform, a value for its weight will be displayed.

Press the TARE key to tare the scale. The weight that was displayed is stored as the tare value and that value is subtracted from the display, leaving zero on the display.

The “TARE” indicator will be on. As product is added only the weight of the product was to be added to the first one. Again only the weight that is added after taring will be displayed.

When the container is removed a negative value will be shown. If the scale was tared just before removing the container this value is the gross weight of the container plus all products that was removed. The zero indicator will also be on because the platform is back to the same condition it was when the ZERO/ENTER key was last pressed.

### WEIGHING A SAMPLE

To determine the weight of a sample first tare an empty container if used,  then place the sample in the container. The display will show the weight and the units of weight currently in use.

### PERCENT WEIGHING

The scale will allow a sample weight to be shown as 100%. Then any other weight placed on the scale will be displayed as a percentage of the original sample.

For example: 700g is placed on the scale and the % key is pressed the display will show 100%. Removing the 700g weight and putting a 200g weight on the scale the display will show 28.57% as 200g is 28.57% of 700g.

Note: the scale may jump by large numbers unexpectedly if small weights are used to set the 100% level. For example if only 23.5g is on a scale with 0.5g increments and the scale is set to 100%, the display will show 100%, however a small change of weight will cause the display to jump to 100.13%.

Pressing the FUNC key will return the scale to weighing.

### PARTS COUNTING

Before beginning, tare the weight of any container that will be used, leaving the empty container on the scale.

When the scale is showing weight, pressing the FUNC key will start the parts counting function.

Place the number of samples on the scale. The number should match the options for parts counting, 10,20,50,100,200,500, or 1000 pieces.

Press the FUNC key to begin. The scale will show “SP 10” asking for a sample size of 10 parts. Change the sample size by pressing the TARE/▲ key. The display will cycle through the options: 10,20,50,100,200,500,1000 and back to 10.

Press the ZERO/ENTER key when the number matches the number of parts used for the sample. As more weight is added the display will show the number of parts (pcs).

Press the % key to display unit weight, Total weight or the count (pcs). Press the FUNC key to return to normal weighing.

Pressing the FUNC key will return the scale to weighing.

### CHECK-WEIGHING

Check-weighing is a procedure to cause an alarm to sound when the weight on the scale meets or exceeds values stored in memory. The memory holds values for a high limit and a low limit. Either limit can be used or both can be used.

Press the Check-W key. The display will show the current High Limit with the left most digit flashing and the HI symbol on the left of the display.

To change the value shown use the %/ and the LIMIT/  to select the digit to change. Then use the TARE/▲ key to increment the flashing digit. When the desired value is shown, press the ZERO/ENTER key to accept the value. If you want to reset the value to zero press the FUNC/C key to clear the value.

After pressing the ZERO/ENTER key the display will then show the Low Limit, The LO symbol will be on the left side of the display.

Enter the low limit in the same way the high limit was entered.

After pressing the ZERO/ENTER key the scale will return to weighing with the Check-weighing function enabled.

When a weight is placed on the scale the arrows will show if the weight is above or below the limits and the beeper will sound as described below.

### BOTH LIMITS SET

The display will show OK and the beeper will sound when the weight is between the limits.

##### LOW LIMIT SET

HIGH LIMIT is set to zero

The display will show OK and the beeper will sound when the weight is less than the Low Limit. Above the Low Limit the display will show HIGH and the beeper will be off.

##### HIGH LIMIT SET

LOW LIMIT is set to zero

The display will show LOW and the beeper will be off when the weight is less than the High Limit. Above the High Limit the display will show OK and the beeper will be on.

##### BOTH LIMITS SET.  LOW IS SET GREATER THAN HIGH

The beeper will never sound and the display will show LOW if the weight is less that the LOW limit, and HIGH if the weight is greater than the Low Limit.

NOTE: The weight must be greater than 20 scale divisions for the check-weighing to operate.

To disable the Check-weighing function enter zero into both limits by pressing the Check-W/ key to recall the current settings. Pressing FUNC/C key when the current limits are shown to clear the settings, then pressing the ZERO/ENTER key to store the zero values.

### ACCUMULATED TOTAL

The scale can be set to accumulate automatically when a weight is added to the scale or manually by pressing the PRINT key. See the PARAMETERS Section for details of selection the method. The accumulation function is only available when weighing. It is disabled during percent weighing or parts counting.

### MANUAL ACCUMULATION

When the scale is set to manual accumulation the weight displayed will be stored in memory when the PRINT key is pressed and the weight is stable.

The display will show “  1” and then the total in memory for 2 seconds before returning to normal. If the optional RS-232 interface is installed the weight will be output to a printer or PC.

Remove the weight, allowing the scale to return to zero and put a second weight on. Press PRINT, the display will show “  2” and then show the new total.

This can continue for up to 99 entries, or until the capacity weight display is exceeded.

To view the totals in memory press the PRINT key when the scale is at zero. The display will show the total number of items “ ××” and the total weight before returning to zero. The totals will also be printed via the RS-232 interface.

To erase the memory press PRINT to view the totals and then press the FUNC/C key to clear the memory.

PARAMETERS

The scale has 6 parameters that can be set by the user. Press the FUNC and PRINT keys at the same time to enter into Parameters Setting model.

The display will show “  ××” The first value shown is the default scale increment value. To select a different increment value press the TARE/▲ key to change the value then press the ZERO/ENTER key.

For example: On a 6kg scale the standard increment is 0.2g, the value can be changed to 0.5g, 1.0g or 2.0g.

The next parameter is control of the backlight. The standard is to have the backlight operate automatically, turning off when the scale is not used. The backlight can be set to be “ ”,“”(Automatic) or “” The maximum battery life is achieved with the backlight turned off.

Press the TARE/▲key to change the value then press the ZERO/ENTER key to confirm.

Print type is next. The default type is “”. The Print type can be set to be “”,“”,“”and“”.

“”: Place on the weight, and the stable mark shows on the left of the display, the scale prints all the data automatically.

“”: Place on the weight, and the stable mark shows on the left of the display, Press the PRINT/ESC key to print the data. When nothing put on the scale, Press the PRINT/ESC key to print the total data.

“”: The RS-232 interface will send the data continuously.

“”: Used by factory only.

Baud Rate setting is next. The default rate is “2400”. The baud rate can be set to be1200, 2400, 4800 ,9600.

Accumulation type is next. The accumulation type can be set to be“”,“”.

Note: The accumulation data should not be print only, the “” type should be choose.

Press the TARE/▲ key to change the value then press the ZERO/ENTER key to confirm.

Select the zero tracking range is next. The value can be set to be “ 0.5”, “ 1”, “ 2”, “ 4”.

Press the TARE/▲ key to change the value then press the ZERO/ENTER key to confirm.

Select the range of stability is next. The value can be set to be “ 0”, “ 1”,“ 2”,“ 3”.

Press the TARE/▲ key to change the value then press the ZERO/ENTER key to confirm.

BATTERY OPERATION

The scale can be operated from the battery if desired. The battery life is approximately 80 hours.

When the battery needs charging a symbol “  ” on the weight display will turn on. The battery should be charged when the symbol is on. The scale will still operate for about 10 hours after which it will automatically switch off to protect the battery.

To charge the battery, simply attach the power supply module to the scale and plug in. The scale does not need to be turned on.

The battery should be charged for 12 hours for full capacity.

There is an LED to indicate the status of battery charging on the right of display. When the scale is plugged into the mains power the internal battery will be charged. If the LED is Green the battery has been   charged. If it is Red the battery is nearly discharged and Yellow indicates the battery is increasing the charge level.

As the battery is used it may fail to hold a full charge. If the battery life becomes unacceptable then contact your distributor.

Note: The battery should be recharged every 3 months if the scale is not used for long time.

RS-232 OUTPUT

The scale can be ordered with an optional RS-232 output.

- Mode EIA-RS 232 C’s UART signal

- Format:

Baud rate: 2400  BPS (varias velocidades configurables)

Data bits:  7    BITS (corregido 8)

Algunos van a 7 bits y el conector trasero. (Mas lejano al teclado)

Parity bit:  1    EVEN

Stop bit:   1    BIT

Code ASCII

Connector: 9 pin socket

Pin3  Output

Pin5  Signal Ground

3. Data Format of Stable transmit, Series transmit, and Transit by pressing PRINT key without accumulation:

HEAD1 (2bytes)

OL-Overload, Under load

ST-Display is stable

US-Display is unstable

HEAD2 (2bytes)

NT-Net weight

GS-Gross weight

Transmit Example:

EX + 0.85kg, when it is stable and net value as:

ST, NT, +0.85 kg   0D    0A

- Transmit Format, when it is in Accumulation model and transmit by pressing PRINT key:

Press the PRINT key

TICKET     No. 01

G           1.0000kg

T           0.0000kg

N           1.0000kg

Press the PRINT key again:

TICKET     No. 02

G           3.0000kg

T           1.0000kg

N           2.0000kg

Press the PRINT key again, when there is nothing on the pan of the scale:

TOTAL NUMBER

OF TICKETS 02

TOTAL

NET     3.0000kg

G=GROSS         T=TARE              N=NET

W=WEIGHT       U=UNIT WEIGHT      Q=QUANTITY

AUTO CALIBRATION

Press the ZERO key and FUNC key at the same time. When the display is flashing and showing the weight value, put on the mass. While the scale is stable, the display will be back to zero automatically.

Please take away the mass while scale is counting backward. The calibration is finished.

By pressing the TARE key the weight value can be selected (1/3 Full weight, 2/3 Full weight and 3/3 Full weight).

RECOVERY OF THE DEFAULT CALIBRATION VALUE FROM MEMORY

Starting up the scale, hold down the TARE key and FUNC key at the same time during self-testing sequence. When the scale begins the self-testing sequence again, and finish self-testing, the scale will come back with a default calibration value.

When the calibration procedure is mistake or parameter is set incorrect, this function should be useful.

ERROR CODES

During the initial power-on testing or during operation it is possible the scale may show an error message. The meaning of the error messages is described below.

If an error message is shown repeat the procedure that caused the message, turning the scale on, calibration or other functions. If the error message still is shown then contact your dealer for further support.

| ERROR CODE | DESCRIPTION | POSSIBLE CAUSES |
| --- | --- | --- |
| E1 | EPROM data lose. |  |
| E2 | A/D count is not correct when turning the scale on | Load cell damaged. Turn on the scale with something on the pan. |
| E3 | The display of percent weighing can not be shown completely. |  |
| E4 | The sample weight is too small. |  |

# Text boxes

+ -